## Slide 1: Neutron Ball Background Testing
Andrea Jedele
Kris Hagel, Alan McIntosh

## Slide 2: Introduction – Neutron Ball
6 segments/tanks
Filled with EJ-335 (pseudocumene and mineral oil) doped with 0.25% wt. nat. Gd
20 PMTs
  4 top
  12 middle
  4 bottom
Beam
direction

## Slide 3: How does the Neutron Ball Work?
Neutrons enter tank
Neutrons interact with liquid
  Thermalized
Thermalized n are captured by Gd
  Emits several gamma rays (1-2 MeV range)
Pseudocumene scintillates
  Photons are collected in PMTs

## Slide 4: Motivation
Reconfiguration of Cave 4
  AGGIE added where Hyperion was
  Hyperion added behind neutron ball
  New beam dump
  Half-wall behind neutron ball removed
  NIMROD remains the same
LLNL modifications
  New liquid scintillator – changed mixture
  New PMTs
Change the neutron ball background rate
  How different is the rate?
  How does this impact our measured neutron multiplicities?

## Slide 5: Cave 4 Configuration
Drawing courtesy of S. Molitor

## Slide 6: Geant simulations
10 MeV neutrons – Simulations courtesy of R. Wada
Effect of half-wall
Effect of Half Wall and Beam Dump Shielding

## Slide 7: Background measurement with Previous configuration
Test ramping up time for K150
Background rate with Sn, Zn and blank
Sn, Zn

## Slide 8: How do we measure neutrons?
Ne-22 at 19 MeV/nuc (K150)
Sn, Zn and blank targets
Phosphor
Si detector as trigger
Scalers
  Counted neutrons in 6 100 μs gates
  Neutron rate in 1st and 2nd 100 μs
  Background rate in subsequent 100 μs gates
Target Wheel
Camera
Si
LED

## Slide 9: Results from “Pulsing beam” – Clamping RF
Measured 4 different times between pulses
3 seconds between beam shut off and ramping up

## Slide 10: Results for target
Beam-on Rate: ~60,000 neutrons/sec.
Neutron Multiplicity
Sn

## Slide 11: Results for Blank target
Blank Target Rate: ~1,000 neutrons/sec.
Neutron Multiplicity

## Slide 12: Phosphor testing
What effect does tuning on a phosphor have on the neutron ball?
Is the neutron ball activated?
  Can we quantify the activation?
  How long is the neutron ball activated?
770 pA on FC02
1 min irradiation
88 nA on FC02
1 min irradiation
12 nA on FC02
1 min irradiation
100 pA on FC02
1 min irradiation
Background

## Slide 13: Extreme case
Event Number
Neutron Multiplicity
95 nA on FC02
Irradiated phosphor for 10 min
Monitored rate after beam turned off
Background

## Slide 14: Next steps
Repeat the same measurement with new beam dump
Test with a higher energy beam

## Slide 15: Thank you!
Alan and Kris
SJY group
Fred, leigh and george

## Slide 16: Pulser trigger measurement
Tested configuration using pulser as trigger vs. Si detector
3 sec
3 sec
Pulser trigger
Si trigger

## Slide 17: How do we measure neutrons?
Si as trigger
Si signal
Pre-Amp
Amplifier
Discriminator
Logic Fan
Target Wheel
Camera
Si
LED

## Slide 18: How do we measure neutrons?
NBL-PMT
Fast Amp
CFD
Linear Fan
Discriminator
Logic Fan
Gate Generator (100μs)
Coincidence
Module
Coincidence
Module
Scalers
  Neutron rate in 1st 100 μs
  Background rate in 2nd 100 μs
Gate Generator (100μs)
Si Discriminator
100μs Delay
Scaler
Scaler
Repeated 5x